“TECHNICAL SPECIFICATION FOR EVALUATION, SUPPLY AND TESTING”

S.T.V.F.C. TC-1/632

Gas chromatograph for gas measurement system

| Compiled by | SERTEC-NORM | IRRERA, ANDREA |
| --- | --- | --- |
| Checked by | SERTEC-NORM | VARESE, ENRICO |
| Checked by | SERTEC-LAB | SALATI, EUGENIO |
| Checked by | PROMA/CMMAT | FASSINO, SILVIO |
| Checked by | PROMA-MATMA | SCOTTO, ENRICO |
| Checked by | MISURA-SVILMIS | GHIA, SERGIO |
| Checked by | HSEQ | BOLZONI, GIANLUCA |
| Approved by | SERTEC | COMAZZI, MARCO |

SUMMARY

1	PRELIMINARY INFORMATION	3

1.1	Cancels and replaces	3

1.2	Reason for issue	3

1.3	Company departments mentioned in the document	3

1.4	Definitions	3

2	PURPOSE AND FIELD OF APPLICATION	3

2.1	Purpose	3

2.2	Field of application	3

3	SPECIFICATIONS	4

4	TECHNICAL CHARACTERISTICS	5

4.1	Sample extraction system	6

4.2	Sample processing system	7

4.3	Gas analysis system with gas chromatograph	7

4.3.1	Gas sample conditioning system	8

4.3.2	Gas chromatograph	9

4.4	Changeover system for service gas cylinders	12

4.5	Housing of gas analysis unit	13

4.6	Electrical power supply system for gas chromatograph	14

4.6.1	Electrical equipment	14

4.6.2	Storage cabinet for electrical equipment and gas chromatograph interface unit	14

4.7	Technical assessment	17

5	Marking	17

6	TYPE TEST	17

7	DOCUMENTATION REQUIRED FROM THE SUPPLIER	18

7.1	Supplier	18

7.2	Documents to be submitted	18

7.3	Test Documentation	18

7.4	Supply Documentation	18

8	TESTS AND INSPECTIONS	19

8.1	Right of the Client	20

9	CONDITIONS OF SUPPLY	20

10	PRODUCT WARRANTY	20

11	REFERENCE STANDARDS	21

12	APPENDICES	22

APPENDIX 1	23

# PRELIMINARY INFORMATION

## Cancels and replaces

This specification cancels and replaces the following Technical Specifications:

- Specification 635 – "Electronic Relative Density Transmitter in Ex-i execution".

## Reason for issue

The update of the Specification 635 and the issue in the form of Technical Specifications for Assessment, Supply and Testing (S.T.V.F.C.), became necessary as a result of:

- new legislative context;

- issuing of new national and European standards (e.g. UNI, CEI, EN, etc.);

- inclusion of Supply and Testing Criteria;

- new company choices.

## Company departments mentioned in the document

- SERTEC-NORM = Standard

- PROMA-CMMAT = Category Manager Materials of Italgas S.p.A.

## Definitions

For definitions of the technical terms used in this specification, please refer to the relevant regulations, a non-exhaustive list of which can be found in Chapter 11.

# PURPOSE AND FIELD OF APPLICATION

## Purpose

This Technical Specification defines the technical and construction characteristics, the accompanying documentation and the procedures for carrying out the technical inspections for the approval for company use of a gas chromatograph to be used in a gas measurement system.

## Field of application

This T.S. concerns the supply of gas analysis system using a gas chromatograph for gas measurement plants (venturi-metric or volumetric measurement plants) installed downstream of the reduction unit, within a WRMP.

It applies, in the case of direct purchase by Italgas Reti, whenever a new PRU covered by this specification must be approved for company use, produced by a new potential Supplier or by an existing Supplier.

# SPECIFICATIONS

- This technical specification must be understood as an integral part of the "ASSESSMENT AND QUALIFICATION SYSTEM FOR SUPPLIERS OF ITALGAS."

- This specification contains the minimum technical requirements demanded by Italgas Reti based on the reference Legislative and Technical Standards.

- It is the Supplier's responsibility to ensure that the products supplied fully comply with the provisions of this technical specification and the applicable standards.

- If the materials in question are supplied by an installer, the installer must document their compliance as required under “Documentation required from the Supplier”.

- It is the Supplier's responsibility, during the product design, production and control process, to verify that there are no construction faults in the finished products (e.g. sharp edges or incorrect directions for use and maintenance), which may jeopardise the health and safety of the Client's workers in relation to the provisions of Italian Legislative Decree no. 81/08, as amended and supplemented.

- For any technical and operational aspects that conflict with or are not mentioned in this technical specification but that nevertheless pertain to the design, installation, maintenance and use of the products themselves, it is the Supplier's responsibility to refer to the regulations in force.

- All the equipment and materials covered by this T.S., with the exception of the gas chromatograph electrical power supply, must be suitable for installation in an area classified ZONE 1 in accordance with CEI EN 60079-10-1.

# TECHNICAL CHARACTERISTICS

The gas chromatograph to be used in a gas measurement system covered by this Technical Specification, is subject to Law 186/68, relating to the realisation and construction of materials, equipment, machinery, installations and electrical and electronic systems “to standard”; it must therefore be suitable for the intended use and ensure, as regards the materials used and their construction, adequate resistance to the environmental and mechanical factors to which they may be exposed during operation.

It is also subject to a whole series of legislative provisions, in relation to the required characteristics, materials and intended use, a non-exhaustive list of which is provided in Chapter 11 of this T.S.

Unless explicitly referred to otherwise, the standards cited are those in the latest edition, including any variants/updates and/or corrigenda published.

The constituent elements of a gas chromatograph for a gas measurement systems are summarised in the table below:

Constituent elements of a gas chromatograph for a gas measurement system

Sample extraction system

Sample processing system

Gas analysis system with gas chromatograph

Changeover system for service gas cylinders

Housing of gas analysis unit

Electrical power supply for gas chromatograph

Table  1: Constituent elements of a gas chromatograph for a gas measurement system

The technical characteristics of the individual elements listed in the table above are indicated below.

## Sample extraction system

The sample extraction system consists of a removable probe and all the elements needed to ensure its safe extraction without depressurising the process line.

The sampling probe must be designed to allow the natural gas sample, taken from the process pipe, to reach the sample processing system as quickly as possible, minimising the time the sample remains in the probe and the disturbance of the natural gas flow at the point in the pipe where the sample is taken.

The probe must be made of seamless stainless steel tube with a diameter of not more than 10 mm in order to achieve a good compromise between sample transport speed, the removal of impurities in the natural gas extracted and insensitivity to vibration, clogging and erosion. The end of the sampling probe should be cut at a 45° angle in order to reduce the possibility of impurities being extracted by the sampling system and should be positioned between 1/3 and 2/3 of the internal diameter of the pipe, as a general rule. The insertion length of the probe must be defined by the Supplier based on the total length from the axis of the pipe to the free surface of the branch connection flange on which the sampling system is to be installed.

The installation of the sampling system must be carried out in process pipes with a nominal diameter of more than 2". For pipes with diameters smaller than those indicated, a pressure intake installed directly on the pipe is sufficient.

It is the responsibility of the Supplier, at the tender stage, to indicate whether appropriate upstream and/or downstream diameters (from the point where the branch connection with flanged attachment is installed) must be complied with.

The sampling system must be suitable for installation in a "hazardous location" (in relation to explosive atmospheres due to the presence of gas), classified as "zone 1" in accordance with CEI EN 60079-10-1 and must be Group IIA (min.) and have a temperature class T3 (min.).

The sample extraction system must have the minimum characteristics listed below:

- Process fluid sampled: natural gas;

- Process fluid temperature: +5°C ÷ 10°C;

- Maximum pressure process fluid: 5 barg;

- Gas density: < 0.8 Kg/m3;

- Process connection: 1" flanged with class 150 according to ASME 16.5;

- Connection to the sample processing system: ¼" NPT-F according to ASME B1.20.1;

- Threaded connections according to ASME B1.20.1.;

- Equipped with a double blocking and vent valve and a probe blocking valve;

- Equipped with a locking bracket that prevents unintentional loosening of the probe or removal of fittings;

- ON-OFF type process fluid isolation valve installed on the probe;

- Construction materials: probe and valves in AISI 316 stainless steel; other materials according to Supplier's standard;

- Assembly: vertical on horizontal pipe;

- Marking with identification of gas flow direction for correct installation.

## Sample processing system

The sample processing system, inserted in the fast loop, must be designed to ensure the availability of fresh, representative gas samples to the gas analysis system with gas chromatograph. The sample processing system must therefore comprise all the equipment needed to provide an initial processing of the natural gas sample, for it to be conditioned and analysed by the gas analysis system with gas chromatograph.

The installation of the equipment constituting the sample processing system must be on an AISI 316 stainless steel plate, suitably sized to contain all the elements needed to carry out conditioning of the extracted gas sample and suitable for wall mounting.

The sample processing system must be suitable for installation in a "hazardous location" (in relation to explosive atmospheres due to the presence of gas), classified as "zone 1" in accordance with CEI EN 60079-10-1 and must be Group IIA (min.) and have a temperature class T3 (min.).

The sample processing system must consist of and not be limited to the following components:

- Two-way ball valve made of AISI 316 stainless steel, ON-OFF type, with double-ring compression connections for 6 mm pipes, enabling isolation of the sample processing system;

- Compact pressure reducer-regulator in stainless steel AISI 316, with ¼" NPT-F threaded connections, calibrated at a pressure that must be within the range of inlet pressures to the gas analysis system with gas chromatograph;

- AISI 316 stainless steel diaphragm filter, with a filtration level of 5 microns (or better), an inlet port and two outlet ports, threaded ¼" NPT-F (one for connection to the gas chromatograph and the other for connection to the fast loop circuit), allowing the elimination of particulate matter and other impurities in the gas sample;

- Bourdon spring pressure gauge made of AISI 316 stainless steel with a 63 mm dial diameter, measuring range suitable for service, with ¼" NPT-M threaded lower process connection;

- Flow meters with structure, wetted parts, float in AISI 304/316 stainless steel with borosilicate glass pipe, fitted with an AISI 316 regulating tap, with measuring range suitable for service;

- Seamless stainless steel tubes with an external diameter of 6 mm and a thickness of 1 mm, double-ring compression fittings for 6 mm tubes, and anything else needed for the pneumatic connections between the various pieces of equipment.

## Gas analysis system with gas chromatograph

The gas analysis system with gas chromatograph consists of two main components:

- The gas sample conditioning system, which must be able to acquire the representative sample of natural gas and transport it to the gas chromatograph;

- The gas chromatograph must be capable of analysing the representative sample of natural gas and communicating the gas parameters to the electronic gas volume conversion device.

The technical characteristics of the individual elements listed above are indicated below.

### Gas sample conditioning system

The gas sample conditioning system, which must be able to acquire the representative sample of natural gas, prevent it from condensing, remove any liquid contaminants and particles and transport it to the gas chromatograph.

The gas sample conditioning system must therefore consist, as a minimum, of the filters to purify the gas from the possible presence of particulate matter, moisture and dust, a flow rate regulator and all the components required to condition the gas sample and ensure that it is acquired by the gas chromatograph.

The natural gas conditioning system must also supply the inlet interface by means of 3 double-ring compression couplings for 6 mm pipes for connection to the sample processing system, the changeover system able to automatically switch between the two service gas cylinders, to the calibration gas cylinder and supply the outlet interface by means of double-ring compression couplings for 6 mm pipes for connection to the atmospheric release pipe, to be defined by the Supplier.

The natural gas sample conditioning system can be integrated into the gas chromatograph itself or directly into the sample processing system. In any case, it is the Supplier's responsibility to include all the elements needed to ensure a correctly conditioned sample of the natural gas in order to avoid damage to the gas chromatograph.

The gas sample conditioning system must be suitable for installation in a "hazardous location" (in relation to explosive atmospheres due to the presence of gas), classified as "zone 1" in accordance with CEI EN 60079-10-1 and must be Group IIA (min.) and have a temperature class T3 (min.).

### Gas chromatograph

The gas chromatograph must be designed for continuous analysis and measurement of natural gas from the sample processing system, conditioned by the gas sample conditioning system. The gas chromatograph must be of a compact type, constructed in such a way as to reduce maintenance time/replacement of components and perform the calculation of the upper calorific value (PCS), lower calorific value (PCI), density (ρ), relative density (d) and Wobbe index (Iw) of the natural gas based on the composition of the gas by mole fraction in accordance with UNI EN ISO 6976.

The gas chromatograph is mainly composed of the gas analysis unit, consisting of the gas chromatographic separation system, the detector(s) and the other components necessary for gas analysis, the processing unit, preferably integrated in the instrument, consisting of the internal electronics, the data storage devices and transmission interfaces for remote communication of the data to external processors and an LCD display with alphanumeric keyboard to carry out maintenance, operation and calibration of the gas chromatograph.

The gas chromatograph must also have OIML R 140 approval as part of a gas measurement system.

The gas chromatograph must be an electrical device suitable for installation in "hazardous locations" (in relation to explosive gas atmospheres), classified as “zone 1” in accordance with CEI EN 60079-10-1 and must belong to Group IIA (min), have a temperature class T3 (min.) and a protection mode (or level) "d" according to CEI EN 60079-0 and CEI EN 60079-1.

The gas chromatograph must have the following minimum characteristics:

- Small size and weight for easy installation in the field;

- Ambient temperature: -20 °C ÷ +60 °C;

- Degree of mechanical protection: IP 65 (min.);

- Equipped with mounting brackets for wall installation;

- Process fluid to be analysed: natural gas.

- Service gas: Helium with guaranteed purity 99.998%;

- Calibration gas: Gas mixture, containing all the components to be analysed, certified by an LAT centre, or EA equivalent;

- Sensor type: Thermo-conductivity sensor;

- Analysis cycle: less than 5 minutes;

- Repeatability: ± 0.025 % CV (0.25 % BTU/scf per 1000 BTU/scf) or better;

- Performance unaffected by external climatic conditions at installation sites;

- Storage of analysis results according to API 21.1;

- Electric power supply: 24 VDC ± 15% from the gas chromatograph electrical power supply, power supply unit with integrated UPS, as described in point 4.6, installed in a "non-hazardous location";

- Electrical connections: no. 3 ISO M32 x 1.5 fitted with adapters from ISO M32 x 1.5 - 1" ISO 7/1;

- Equipped with earth terminal with single-core cable for fixed installation with a minimum cross-section of 6 mm2;

- Management of reduced consumption of service and calibration gas;

- Automatic restart after power failure with predetermined cyclic sequence;

- Maintenance of operating programme for at least 30 days in the event of a power failure;

- Analysis of the gas components, i.e. methane (C1), ethane (C2), propane (C3), isobutane (iC4), n-butane (nC4), isopentane (iC5), n-pentane (nC5), hexanes and higher (C6+), nitrogen (N2), carbon dioxide (CO2) and determination of the concentrations of the components of the individual analyses, normalised to 100 and rounded to three decimal places (for rounding, the value of methane is calculated by difference from 100);

- Indication of the composition of the gas normalised to 100% with indication of the non-normalised total; the normalised composition includes the percentage of helium, which it must be possible to enter either as a fixed value or as an automatically calculated value correlated to the methane content;

- Recognition and indication of instrumental faults;

- Linearity of response throughout the range of permissible concentrations for individual components;

- Automatic calibration at programmable frequency;

- Complete with serial ports configurable as RS232/RS485 implementing the MODBUS RTU/ASCII protocol for interfacing with the computerised gas volume measurement system, especially the electronic gas volume conversion device;

- Complete with an insulated, programmable digital input, capable of accepting an ON-OFF contact from the changeover system for service gas cylinders: the Gas Chromatograph Service Gas Cylinder Empty (BS) alarm acquired by the gas chromatograph must be transmitted via serial to the measuring cabinet;

- Equipped with a programmable digital output, in the form of a voltage-free, normally open (NO) contact, in fail-safe mode, with a capacity of 1 A (max.) at 24 VDC for transmission of the Gas Analysis System Anomaly Alarm with Gas Chromatograph (AAG)  ;

- Complete with a serial or Ethernet communication port to which a gas chromatograph interface unit can be connected, i.e. an electronic device, to be installed in a "non-hazardous location" and inside the cabinet containing the gas chromatograph electrical supply system, by means of which an external portable electronic device can be connected that can download and display the measured values locally;

- Complete with alphanumeric keypad and LCD display for management, calibration and maintenance operations and display of measured variables and processed values on demand.

## Changeover system for service gas cylinders

The changeover system for the service gas cylinders must be designed to ensure a continuous gas flow for the gas analysis system with gas chromatograph and automatic switching between the two service gas cylinders when the pressure inside the cylinder drawn from is no longer sufficient to deliver the gas.

The equipment constituting the changeover system for the service gas cylinders must be built on an AISI 316 stainless steel plate, suitably sized to contain all the necessary elements and suitable for wall mounting.

The changeover system for service gas cylinders must be suitable for installation in a "hazardous location" (in relation to explosive atmospheres due to the presence of gas), classified as "zone 1" in accordance with CEI EN 60079-10-1 and must be Group IIA (min.) and have a temperature class T3 (min.).

The changeover system for service gas cylinders must consist of and not be limited to the following components:

- A compact pressure reducer-regulator in AISI 316 stainless steel, with threaded ¼" NPT-F connections, which can be calibrated manually to the pressure necessarily within the range of inlet pressures to the gas analysis system with gas chromatograph. It must also be possible to install a Bourdon spring pressure gauge in AISI 316 stainless steel, with a measuring range suitable for service and threaded ¼" NPT-M (min.) lower process connection, which allows the pressure to be measured downstream of the changeover system for the service gas cylinders;

- Switching regulator, which performs the function of automatic changeover between the two service gas cylinders allowing the safe manual replacement of the empty cylinder, on which it is possible to install 2 Bourdon spring pressure gauges made of AISI 316 stainless steel, with a measuring range suitable for service and ¼" NPT-M (min.) threaded lower process connection, equipped with electrical contacts, which allow the downstream pressure of each service gas cylinder to be measured and an empty cylinder alarm to be transmitted;

- No. 1 junction box in Ex d execution according to CEI EN 60079-0 and CEI EN 60079-1 standards, equipped with a terminal board on which the two alarm contacts coming from the two pressure gauges dedicated to the pressure measurement of the respective service cylinders can be connected in series.

## Housing of gas analysis unit

The gas analysis unit housing is the self-supporting structure that must be designed and constructed for the installation and support of the following components:

- Sample processing system;

- Gas analysis system with gas chromatograph;

- Changeover system for service gas cylinders;

- No. 2 service gas cylinders and no. 1 calibration gas cylinder.

The entire structure is to be made of quality steel and the steel materials (i.e. profiles or pipes, bars, plates, etc.) are to be supplied assembled and welded together so as to create a self-supporting structure of small and compact dimensions, which is then to be fixed to the floor.

The sample processing system and the gas analysis system with gas chromatograph must be installed on the front of the self-supporting structure, while on the back side 2 service gas cylinders, typically of 50 L each containing helium gas, the calibration gas cylinder, typically of 10 L, and an AISI 316 stainless steel pressure reducer-regulator with ¼" NPT-F threaded connections must be installed. It must be possible to calibrate the latter manually to the pressure necessarily within the range of inlet pressures to the gas analysis system with gas chromatograph, on which it is possible to install 2 Bourdon spring pressure gauges in AISI 316 stainless steel, with a measuring range suitable for service and threaded ¼" NPT-M (min.) lower process connection, which allows the pressure to be measured downstream of the pressure reducer-regulator. The calibration gas cylinder must be suitably heated to prevent condensation, in accordance with the requirements of UNI 9167.

The entire structure must be easily accessible in all its parts in order to allow maintenance operations of the individual components installed in the structure.

The cylinders must be secured in such a way that they can be easily replaced if necessary.

The structure must be equipped with an earthing bar to connect the individual metal enclosures by means of 6 mm2 (min.) PVC- or rubber-insulated copper cables to the main earthing system.

A seamless stainless steel pipe with an external diameter of ¾" (min.) must be installed on the structure to which all equipment vents are connected, on one end of which a vent valve is installed.

## Electrical power supply system for gas chromatograph

The gas chromatograph electrical power supply system consists of two main components:

- Electrical equipment;

- Storage cabinet for electrical equipment and gas chromatograph interface unit.

The technical characteristics of the individual elements are listed below.

### Electrical equipment

The gas chromatograph electrical power supply system must have the following features and components:

- suitable for mains supply voltage 230 VAC ±15% at 50 Hz ±5%;

- electrical protections sized and aligned on all input and output circuits;

- any additional overvoltage protections and mains filters, if necessary, considering that upstream SPDs of type T1 and T2 are installed as per T.S. IE-1/441;

- 24 VDC power supply system suitable, in addition, to guarantee an autonomy of more than 20 hours of the gas analysis system in the event of a mains power failure;

- filter to prevent the emission of radio interference;

- pre-alarm, alarm and total lockout thresholds (with manual reset) connected to back-up batteries with "self-test" circuit on battery function with voltage-free contacts for remote re-transmission of generated signals;

- complete with voltmeter and amperometer on the low voltage circuit installed on the front panel.

### Storage cabinet for electrical equipment and gas chromatograph interface unit

The cabinet containing the electrical equipment, described in point 4.6.1, must comply with the standards CEI EN 61439-1 and CEI 23-49, be suitable for wall mounting, and therefore be equipped with all necessary mounting accessories.

The electrical panel must be installed in a "non-hazardous location" (in relation to explosive atmospheres due to the presence of gas), in accordance with CEI EN 60079-10-1, have a mechanical protection rating of not less than IP 44, in accordance with CEI EN 60529, and a strength rating of not less than IK 09, in accordance with CEI EN 62262.

The storage cabinet must have the following mechanical characteristics:

- small dimensions and weight compatible with the installation requirements based on the type of system requested in the order, aimed at minimising installation, commissioning and maintenance costs;

- it must be self-supporting and composed of press-moulded and welded, pickled, passivated sheet steel panels with a minimum thickness of 12/10 mm in light grey RAL 7035 epoxy paint with a "fine textured" finish;

- inner surface treated with anti-condensation material;

- inspectable from the front by means of a transparent door made of safety glass, i.e. tempered glass in accordance with UNI EN 12150, fitted with hinges permitting 180° opening and equipped with closing handles, hinged on both the right and left sides, preferably double to allow full opening even in confined spaces;

- internal counter-door, suitably hinged to allow access to internal parts for normal maintenance, on which the amperometer, voltmeter and RS232 port of the gas chromatograph interface unit are present, located on the transparent door;

- cable entry possible from both the top and bottom of the panel through easily drilled panels complete with cable glands;

- support plate, fixed to the rear wall, for terminal blocks;

- strips of interconnecting cables between switches and equipment located on the panels and terminal blocks, contained in special self-extinguishing sheaths;

- cable ducts, DIN profiles and related terminal blocks mounted on plates;

- internal document pocket for various wiring diagrams for maintenance, etc.;

- small (unsealed) compartment for storing "immediate repair" spares, i.e. fuses, arresters, etc.

The electrical panel must have the following accessories and components:

- warning signs appropriately placed and compliant with Legislative Decree no. 81/08 as amended, Title V;

- mechanical and electrical protection complying with the CEI EN 60947 series of standards for each piece of equipment in the system that is to be individually protected and signalled;

- input/output terminal blocks, and wiring in general, appropriately marked, protected against accidental contact, easily accessible and placed in accordance with the standards of the CEI EN 61439 series and Law 186/68;

- adequately ventilated compartment by means of slits;

- earthing bar (equipotential node) to which all normally non-energised parts of the electrical equipment and carpentry will be connected.

The wiring must use flexible wire, with very low emission of fumes and toxic gases, not propagating fire and with Uo/U rated voltages not exceeding 0.6/1 kV according to CEI 20-38 and must be divided into the following categories:

- Rated voltages at 24 VDC;

- Rated voltages at 230 VAC.

The markings of the conductors connected to terminal blocks must comply with CEI EN 60445.

Each end of the conductor will be marked using both the symbol of the terminal to which it is connected and that of the terminal to which the other end is connected.

The marking of each end will be completed by the symbol of the terminal block, relay or equipment, etc. to which such end is connected.

The terminals will be marked with acronyms, progressive numbers and shown on the drawings.

In particular:

- the cables should, where possible, have a different colour according to the voltage level;

- the conductor cross-sections must be suitable for the intended load;

- the conductor terminals must be fitted with a suitable cable lug and numbered with marking protected with transparent tube;

- all cables will be laid in plastic ducts, and space must be left in such ducts to add 50% more "spare" cable in the future;

- splicing of cables will not be permitted;

- cables and ducts must not interfere with maintenance and equipment testing;

- interconnecting cables fitted with appropriate connectors to allow easy connection between terminal blocks and components inside the electrical panel;

- the cable shields should be connected to the equipotential bonding bar inside the panel, keeping the connection as short as possible. The shield should be earthed at one end only;

- the earthing conductors and terminals must be green/yellow;

- adequate protection must be provided, where necessary, on the terminals and normally live components inside the panel in order to prevent accidental contact by staff.

An earthing bar (equipotential node) must be provided on the panel.

The closing doors and revolving frame must be connected by means of flexible cables of a suitable cross-section to the earthing bar (equipotential node).

All normally non-energised parts of the electrical equipment must be connected by means of a flexible cable to the earthing bar (equipotential node).

## Technical assessment

In relation to Legislative Decree no. 81/08, as amended and supplemented, Italgas Reti reserves the right, in its evaluation of the use of the proposed products, to consider the operational consequences deriving from the conditions of use of the equipment as prescribed by the Supplier in the "Use and Maintenance Manual".

# Marking

The products covered by this T.S. must be provided with the markings required by the EU Directives and by the relevant applicable technical standards, considering that the affixing of the CE marking, together with any other markings that may be used, is subject to the provisions of Decision no. 768/2008/EC of 9 July 2008 on a common framework for the marketing of products, implemented by the legislative decrees transposing the applicable product directives, a non-exhaustive list of which is given under "Legal references" in this T.S.

They must therefore bear the conformity marking , affixed in the manner provided for by the legal provisions in force, together with any other markings provided for by the same directives (e.g. "specific marking of explosion protection”  in accordance with the ATEX Directive, “marking relating to the disposal of electrical and electronic equipment"  in accordance with the WEEE Directive, etc.).

# TYPE TEST

In the design stage of the gas chromatograph for a gas measurement plant, the "Type Tests" required by the Reference Standards and/or the Supplier's design standards and by the standards referred to therein, for verifying the intended end use of the product must have been performed.

# DOCUMENTATION REQUIRED FROM THE SUPPLIER

## Supplier

- The potential Supplier must be qualified as an Italgas Supplier.

## Documents to be submitted

The potential Supplier must submit the following documentation to the PROMA-CMMAT Department at the tender stage:

- “Declaration of Conformity” with this technical specification and any specific documents mentioned in it (legislative provisions, applied standards, etc.) to which the Supplier refers for the design and construction of the product, in accordance with the requirements of UNI CEI EN ISO/IEC 17050-1 (see the example in APPENDIX 1 to this technical specification).

- Whatever else the potential Supplier considers useful to better qualify the proposed product.

## Test Documentation

The Supplier must issue a “type 3.1” Inspection Certificate in accordance with UNI EN 10204, which states that the products provided are compliant with the requirements of these specifications and which provides the test results.

Registration of the full production/test documentation completed by the Supplier must be kept (on the basis of legal provisions) for 10 years under its responsibility, and must be made available whenever requested by the Client.

## Supply Documentation

The following documentation is required from the Supplier when supplying the gas chromatograph for a gas measurement system:

- a statement certifying ownership to market the product;

- "EU Declaration of Conformity" in accordance with legislative decrees implementing the Directives applicable to the product (BT, PED, ATEX, EMC, RAEE, RoHS II Directives etc.);

- EU Type Examination Certificate (Directive 2014/34/EU);

- certificate confirming compliance with OIML R140 as an integral part of a measurement system;

- calibration certificate;

- test report including all tests performed in the factory and results of the same;

- copy of the "Installation, Use and Maintenance Manual" (in Italian) for the correct use of the product;

- documents compliant with Legislative Decree no. 286 of 21 November 2005, as amended;

- any other technical documents and/or drawings requested by the relevant Italgas Reti departments.

# TESTS AND INSPECTIONS

Italgas Reti reserves the right:

- to be present with its own Inspectors and/or Inspection Companies appointed by the same to carry out tests/checks at the production plant and at the Supplier's care and expense, to check the constructive and functional compliance of the materials in accordance with the specifications and/or reference standards.

- to verify, upon receipt and/or under actual conditions of use, the perfect compliance of the goods with the functional technical characteristics expressly indicated in the Contract.

In any case, before use in the field, the characterisation of the equipment must be carried out by the Supplier, consisting of verification of repeatability and accuracy as indicated in the UNI 9167 standard.

The repeatability check, according to Table  2, using as test gas the self-calibrating gas mixture, which must be certified by a LAT Centre, or EA equivalent, consists of performing at least 7 consecutive analyses of a gas sample containing all the components to be determined, discarding the first two analyses.

| C1 – C2 | 0.1 % molar |
| --- | --- |
| C3 – N2 – CO2 | 0.05 % molar |
| PCS – PCI | 50 kJ/Sm3 |
| Relative density | 0.001 |
| Zs (at reference conditions) | 0.001 |

Table  2: Verification of repeatability

The verification of accuracy for PCS, PCI, Relative Density, Z, CO2 and N2 using two samples of test gas, containing all the components to be determined with PCS between 37.3 MJ/Sm3 ÷ 38.1 MJ/Sm3 and 38.9 MJ/Sm3 ÷ 40.2 MJ/Sm3, certified by a LAT Centre, or equivalent in EA, consists of performing 5 analyses, discarding the first two for each test sample. On the last three analyses, the average composition and the relevant chemical-physical parameters are calculated, verifying that the relative error calculated by comparison with the certificate of analysis of the test gas is within the limits given below.

| PCS – PCI | 0.5% |
| --- | --- |
| Relative density | 0.5% |
| Zs (at reference conditions) | 0.1% |
| χCO2 | 0.1 χCO2 |
| χN2 | 0.1 χN2 |

Table  3: Verification of accuracy

It is understood that the favourable outcome of our inspection and checking of the materials shall not release you in any way from the responsibilities and obligations arising from each individual contract.

## Right of the Client

In order to verify the conformity of the proposed products with the contents of this technical specification and the cited standards to which reference is made, it is Italgas Reti's right, for each purchase order, to request from the Supplier:

- the production of further relevant technical and/or graphic documentation with the indicated characteristics;

- the performance of any additional laboratory and/or field tests in relation to the required product characteristics.

# CONDITIONS OF SUPPLY

Supplier's Standard in compliance with Legislative Decree no. 152 of 3 April 2006 as amended and supplemented.

# PRODUCT WARRANTY

The metrological performance and construction characteristics of the electronic gas volume conversion device under normal use must be guaranteed for 3 years; the warranty applies if the deviation in performance is attributable to a defect in the electronic gas volume conversion device and ex works.

All electronic gas volume conversion devices that for any reason (periodic, random) have been inspected with a negative outcome must be replaced under warranty, on the basis of the warranty terms set out above.

# REFERENCE STANDARDS

- Legislative Decree no. 86 of 19 May 2016 - "Implementation of Directive 2014/35/EU on the harmonisation of the laws of Member States relating to the making available on the market of electrical equipment designed for use within certain voltage limits";

- Legislative Decree no. 85 of 19 May 2016 – "Implementation of Directive 2014/34/EU on the harmonisation of the laws of the Member States relating to equipment and protective systems intended for use in potentially explosive atmospheres";

- Legislative Decree no. 194 of 06 November 2007, as amended. - "Implementation of Directive 2004/108/EC on electromagnetic compatibility, and Directive 2014/30/EU of 26 February 2014 on the harmonisation of the laws of the Member States relating to electromagnetic compatibility (recast) providing for its repeal";

- Legislative Decree no. 49 of 14 March 2014, as amended. - 'Implementation of Directive 2012/19/EU on waste electrical and electronic equipment (WEEE)';

- Legislative Decree no. 27 of 04 March 2014, as amended. - 'Implementation of Directive 2011/65/EU on the restriction of the use of certain hazardous substances in electrical and electronic equipment';

- Legislative Decree no. 93 of 25 February 2000, as amended. - "Implementation of Directive 97/23/EC on pressure equipment and Directive 2014/68/EU on the harmonisation of the laws of the Member States relating to the making available on the market of pressure equipment (recast), providing for its repeal";

- Legislative Decree no. 81 of 09 April 2008, as amended. - “Implementation of Article 1 of Law no. 123 of 3 August 2007 on the protection of health and safety in the workplace”;

- Legislative Decree no. 152 of 03 April 2006, as amended. - “Environmental standards”;

- CEI EN 60079-0 (CEI 31-70) - “Explosive atmospheres - Part 0: Equipment - General Requirements";

- Law no. 186 of 1 March 1968 - "Provisions concerning the production of materials, equipment, machinery, electrical and electronic installations and systems";

- CEI EN 60079-1 (CEI 31-58) “Explosive atmospheres - Part 1: Equipment protected by flameproof' enclosures ‘d’;

- CEI EN 60079-10-1 (CEI 31-87) “Explosive atmospheres - Part 10-1: Classification of areas - Explosive gas atmospheres';

- CEI EN 60529 (IEC 70-1) - 'Degrees of protection provided by enclosures (IP Code)';

- CEI EN 62262 (CEI 70-4) – "Degrees of protection provided by enclosures of electrical equipment against external mechanical impacts (IK Code)";

- CEI EN 61439-1 (CEI 17-113) – "Low-voltage switchgear and controlgear assemblies (LV panels)";

- CEI EN 62040 – "Uninterruptible power systems (UPS)";

- CEI EN 60445 (CEI 16-2) – "Basic and safety principles for man-machine interface, marking and identification - Identification of equipment terminals and conductor ends";

- CEI EN 60947 – "Low Voltage Equipment";

- CEI 20-38 – "Halogen-free, rubber-insulated, fire retardant cables for rated voltages U0/U not exceeding 0.6/1 kV";

- CEI 23-49 – "Enclosures for appliances for household and similar fixed electrical installations - Part 2: Special requirements for enclosures intended to contain protective devices and equipment which in normal use dissipate non-negligible power";

- CEI EN 61000 –Electromagnetic compatibility (EMC)";

- CEI EN 60068 – "Environmental testing";

- CEI EN 61326-1 – "Electrical equipment for measurement, control and laboratory use - EMC requirements - Part 1: General Requirements";

- UNI EN 12150 – "Glass in building - Thermally toughened soda lime silicate safety glass";

- UNI EN ISO 6976 – "Natural gas - Calculation of calorific value, density, relative density and Wobbe index from composition";

- UNI EN ISO 13443 – "Natural gas - Standard reference conditions";

- UNI CEI EN ISO/IEC 17050-1 - “Conformity assessment - Supplier’s declaration of conformity. Part 1: General requirements”;

- UNI 9167 – "Natural gas receiving, first reduction and measuring systems - Design, construction and testing";

- UNI 9571- 2 – Gas infrastructures. Gas pressure control and metering stations connected with transportation systems. Part 2: Surveillance of measurement systems;

- OIML R140 – "Measuring systems for gaseous fuel".

# APPENDICES

- APPENDIX 1 -	“Sample declaration of conformity with this technical specification”

## APPENDIX 1

Supplier's Declaration of Conformity

(in accordance with the UNI CEI EN lSO/IEC 17050-1 standard)

(1)	No. ...............................................

(2)	Name of the issuer: ..............................................................................................................................................................................

Address of the issuer: .........................................................................................................................................................................

(3)	Subject of the declaration: ...................................................................................................................................................................

...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

The subject of the above declaration complies with the requirements of the following documents:

Document nos.	Title	Edition/Date of issue

(4)	S.T.V.F.C. TC-1/632	Gas Chromatograph for gas measurement system	Edition 1 of 15/06/2017

Additional information:

(5) 	...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

Signed for and on behalf of:

.............................................................

(Place and date of issue)

(6) 	............................................................. 			......................................................................................................................

(Name and position) 						(Signature or equivalent mark authorised by the issuer)

KEY:

- The Declaration must be issued on headed paper or, if visible, on applied stamps and be uniquely identified (e.g. by a log number).

- The person issuing the Declaration (Supplier) must be unequivocally specified.

- List all the products for which approval for company use is requested (name, type, model), stating any codes assigned by the Supplier itself and, where existing, the relevant material codes assigned by Italgas.

- Include references to the Italgas Reti documents for the proposed products.

- Indicate any deviations from the reference documentation and any other information deemed useful by the Supplier for product evaluation purposes (see note 6) of appendix A of UNI CEI EN ISO/IEC 17050-1).

- Include the handwritten signature of the person authorised to issue the Declaration, stating their full Name, Surname and Position within the Organisation.

Supplier means the party responsible for the design and construction of the products or its legal representative.

The Client is entitled to request assembly different from that indicated, in the event of special design requirements, which must be communicated to the Supplier at the tender stage. In any case, it is the Supplier's responsibility to inform the Client if there are any limitations in the installation of the sampling extraction probe and to make a suitable proposal.

The pressure reducer-regulator should be installed as close as possible to the sample extraction system, while taking into account the time for gas chromatograph analysis, in order to reduce delays.

For the operation and maintenance of the gas chromatograph installed in a WRMP, reference should also be made to the requirements of UNI 9571-2 in addition to those defined in the Supplier's Instruction and Maintenance Manual.

If the processing unit with display is not integrated in the gas chromatograph and its installation in a "non-hazardous location" is required, the Supplier must provide all the necessary information for the electrical connections between the gas chromatograph and the processing unit. In addition, this proposal must be accepted by the Client at the tender stage.

If the gas chromatograph proposed by the Supplier operates in a lower temperature range than the one mentioned, then the Supplier must include in the scope of supply a box with an air conditioning system to ensure optimal environmental conditions for the gas chromatograph inside the box and use electronic components suitable for installation in an area classified as "zone 1", in particular Ex d IIA T3 (min) equipment. This proposal/solution must be evaluated and confirmed at the tender stage by the Client.

The Supplier must also provide ATEX-certified caps for unused inlets.

The Supplier must indicate the actual consumption of service gas and calibration gas at the tender stage.

In the event that the gas chromatograph does not have the contact available, the Gas Chromatograph Service Gas Cylinder Empty (BS) alarm can be acquired by the RTU, subject to verification of signal acquisition availability in the Remote Control Cabinet.

The Supplier must check/confirm the possibility of sending the Gas Analysis System Anomaly Alarm with Gas Chromatograph (AAC) via serial to the measuring cabinet as an alternative.

The Supplier must indicate at the tender stage the number of available ports and the specific function/configuration of the same.

The supply of the service and calibration gas cylinders must cover the requirements for the start-up of use of the instrument and ensure its operation for at least a period of about six months.

Near the entry and exit points of the cables to and from the field, suitable profiles must be provided for anchoring the cables.

Appropriate cable ducts must be provided for cables coming from/to the field, dedicated exclusively to the same.