“TECHNICAL SPECIFICATION FOR EVALUATION, SUPPLY AND TESTING”

S.T.V.F.C. 9993302

Portable methane gas and LPG detectors

| Compiled by | SERTEC-NORM | IRRERA, ANDREA |
| --- | --- | --- |
| Checked by | SERTEC-NORM | VARESE, ENRICO |
| Checked by | SERTEC-LAB | SALATI, EUGENIO |
| Checked by | PROMA/CMMAT | FASSINO, SILVIO |
| Checked by | PROMA-MATMA | SCOTTO, ENRICO |
| Checked by | HSEQ | BOLZONI, GIANLUCA |
| Approved by | SERTEC | COMAZZI, MARCO |

SUMMARY

1	Preliminary information	3

1.1	Cancels and replaces	3

1.2	Reason for issue	3

1.3	Company departments mentioned in the document	3

1.4	Definitions	3

2	PURPOSE AND FIELD OF APPLICATION	3

2.1	Purpose	3

2.2	Field of application	3

3	SPECIFICATIONS	4

4	General technical characteristics	4

4.1	Description and characteristics of portable methane gas detectors	5

4.1.1	Technical characteristics relevant to instruments without a laser pointer	5

4.1.2	Technical characteristics relevant to laser pointer instruments	6

4.1.3	Construction characteristics	7

4.1.4	Accessories	7

4.2	Description and characteristics of portable LPG detectors	9

4.2.1	Technical characteristics	9

4.2.2	Construction characteristics	9

4.2.3	Accessories	9

4.3	Technical assessment	9

5	MARKING	9

6	TYPE TEST	9

7	DOCUMENTATION REQUIRED FROM THE SUPPLIER	10

7.1	Supplier	10

7.2	Documentation for the company approval process	10

7.3	Test Documentation	11

7.4	Supply Documentation	11

8	TESTS AND INSPECTIONS	11

8.1	Test sample	11

8.2	Type of tests	11

8.3	Right of the client	12

9	CONDITIONS OF SUPPLY	12

10	REFERENCE STANDARDS	12

11	APPENDICES	13

APPENDIX 1	14

# Preliminary information

## Cancels and replaces

This specification cancels and replaces technical specifications 9993302 edition 1 of 30 December 2011.

## Reason for issue

The update of the Technical Specification (T.S.) became necessary as a result of:

- new legislative context;

- issuing of new national and European standards (e.g. UNI, CEI, EN, etc.);

- updating of Supply and Testing Criteria.

## Company departments mentioned in the document

- SERTEC-NORM = Standard

- SERTEC-LAB = Laboratory.

## Definitions

For definitions of the technical terms used in this specification, please refer to the relevant regulations, a non-exhaustive list of which can be found in Chapter 10.

# PURPOSE AND FIELD OF APPLICATION

## Purpose

To define the technical and constructional characteristics, the accompanying documentation and the procedures for carrying out the technical inspections for the approval for company use and the criteria for the supply and testing of portable methane and/or liquefied petroleum gas (LPG) detectors.

## Field of application

It applies, in the case of direct purchase by Italgas Reti, whenever a material covered by this T.S. must be approved for company use, produced by a new potential Supplier  or by an existing Supplier.

In particular, this T.S. applies to portable equipment enabling the continuous detection and measurement of airborne concentrations of:

- methane gas;

- liquefied petroleum gas (LPG),

to be used for the following activities:

- pre-location of gas leaks from the underground network;

- location of gas leaks from the underground network;

- other routine or emergency activities related to methane gas and LPG detection in open, closed or poorly ventilated environments.

Portable equipment using the following technologies falls within the scope of this T.S.:

- semiconductor sensors;

- catalytic sensors;

- thermal conductivity sensors;

- portable equipment using flame ionisation technology (for pre-location activities only);

- laser sensors (with sample suction in a measuring chamber inside the equipment; or laser pointer with infrared measuring beam and visible pointer beam).

# SPECIFICATIONS

- This technical specification must be understood as an integral part of the "ASSESSMENT AND QUALIFICATION SYSTEM FOR SUPPLIERS OF ITALGAS.”

- This specification contains the minimum technical requirements demanded by Italgas Reti based on the reference legislative and technical standards.

- It is the Supplier's responsibility to ensure that the products supplied fully comply with the provisions of this technical specification and the applicable standards.

- It is the manufacturer’s responsibility, during the product design, production and inspection process, to verify that there are no construction faults in the finished products (e.g. sharp edges, incorrect directions for use and maintenance, etc.), which may jeopardise the health and safety of the Client's workers in relation to the provisions of Legislative Decree no. 81/08, as amended.

- For all technical and operational aspects that conflict with or are not mentioned in this technical specification, but nevertheless pertain to the design, maintenance and use of the products themselves, it is the Supplier's responsibility to refer to the regulations in force.

# General technical characteristics

All equipment comprising the detector must be designed and constructed in application, as far as relevant, of the requirements of:

- Law no. 186 of 1 March 1968 (hereinafter referred to as Law 186/68);

- Legislative Decree No. 194 of 6 November 2007 as subsequently amended and supplemented (hereinafter referred to as Legislative Decree 194/07);

- Legislative Decree No. 81 of 9 April 2008 as subsequently amended and supplemented (hereinafter referred to as Legislative Decree 81/08);

- Legislative Decree no. 85 of 19 May 2016 (hereinafter Legislative Decree 85/16).

In particular, it must meet the requirements of CEI EN 61010-1 and, where appropriate, the standards of the same series with regard to the safety of electrical equipment for measurement, control and laboratory use.

Specifically, the detectors in question must be manufactured in accordance with the requirements of standard CEI EN 60079-29-1 with regard to construction, testing and required performance, with the exception of portable detectors using laser technology.

With regard to laser pointer detectors, these must also comply with CEI EN 60825-1.

The portable detectors covered by this T.S. divide into two classes:

- portable methane gas detectors;

- portable LPG detectors.

Portable gas detectors with dual measurement functionality for both natural gas and LPG may comply with the following specification.

## Description and characteristics of portable methane gas detectors

The device must be robust, light and have the following characteristics.

### Technical characteristics relevant to instruments without a laser pointer

Portable detectors using flame ionisation may only be used for leak pre-location activities.

Equipment with a measuring range of 0 to 100% LEL must be supplied and certified with an LEL/LFL value of 4.4% in volume for methane gas.

- Measuring ranges:

  - for leak pre-location activities:

- 0  10,000 p.p.m.v. (maximum value, measuring ranges with a lower full scale can also be used, up to a minimum value of 1,000 p.p.m.v.).

If the instrument is equipped with several measuring ranges in p.p.m.v., there may also be measuring ranges with a lower full scale.

- 0  100 % LEL (Lower Explosive Limit, optional measuring range).

- 0  100 % by volume (optional measuring range).

  - for other activities:

- 0  10,000 p.p.m.v. (maximum value, measuring ranges with a lower full scale can also be used, up to a minimum value of 1,000 p.p.m.v.).

- 0  100 % LEL (Lower Explosive Limit).

- 0  100 % by volume.

- Resolution:

  - for leak pre-location activities:

- 0  100 p.p.m.v.: resolution ≤ 1 p.p.m.v;

- 100  1,000 p.p.m.v.: resolution ≤ 10 p.p.m.v;

- over 1,000 p.p.m.v.: resolution ≤ 100 p.p.m.v.

  - for other activities:

- the resolution must be ≤ 1% of the relevant full scale value.

- Measurement errors:

  - for leak pre-location activities:

- 0  100 p.p.m.v.: measurement error ≤ 10 p.p.m.v;

- 100  1,000 p.p.m.v.: measurement error ≤ 100 p.p.m.v;

- over 1,000 p.p.m.v.: measurement error ≤ 1,000 p.p.m.v.

  - for other activities:

- as regards LEL and volume scales, the error must be ≤ 10% of the respective full scale value;

- with regard to the p.p.m.v. scale, the following applies:

- 0  1,000 p.p.m.v.: measurement error ≤ 150 p.p.m.v;

- 1,000  10,000 p.p.m.v. (for instruments with this full scale): measurement error ≤ 1,000 p.p.m.v.

- Withdrawal pump flow rate:

  - for leak pre-location activities:

- ≥ 50 l/h;

  - for other activities:

- ≥ 30 l/h.

- Environmental conditions of use:

  - for leak pre-localization activities, performed with flame ionisation detectors:

- temperature range: -5°C to 40°C;

- relative humidity range: 20% to 90%;

  - for leak pre-localisation activities, performed with non-flame ionisation detectors:

- temperature range: -15°C to +40°C;

- relative humidity range: 5% to 80%;

  - for other activities, performed with non-flame ionisation detectors:

- temperature range: -20°C to +40°C;

- relative humidity range: 0% to 90%.

- Degree of protection: at least IP54.

- Minimum requirements regarding ATEX certification:

  - equipment group: at least IIA;

  - category: at least 2G;

  - temperature class: at least T3;

  - protection mode: single or combined, from among those permitted for category 2G (preferably "i" - intrinsically safe, "d" - flameproof enclosures, "e" - increased safety, "m" - encapsulation).

The relevant marking must comply with the requirements of Legislative Decree 85/16 and standard CEI EN 60079-0.

- Equipped with rechargeable batteries allowing a total autonomy of at least eight hours of continuous operation, or non-rechargeable batteries with an autonomy of at least thirty hours. It must be possible to check the charging status of the batteries.

- Calibration must be possible in less than half an hour by SERTEC-LAB personnel using calibration equipment/benches.

### Technical characteristics relevant to laser pointer instruments

- Measuring range: 0 ÷ 50,000 p.p.m.×m.

- Sensitivity: 10 p.p.m.×m.

- Measurement error: 100 p.p.m.×m (100 ~ 1,000 p.p.m.×m).

- Detection distance: maximum 30 m.

- Environmental conditions of use:

  - temperature range: -15°C to +50°C;

  - relative humidity range: 30% to 90%.

- Measuring laser beam power: Class 1.

- Laser pointer beam power: Class 2.

- Degree of protection: at least IP54.

- Minimum requirements regarding ATEX certification:

  - equipment group: at least IIA;

  - category: at least 2G;

  - temperature class: at least T1;

  - protection mode: single or combined, from among those permitted for category 2G (preferably "i" - intrinsically safe, "d" - flameproof enclosures, "e" - increased safety, "m" - encapsulation).

The relevant marking must comply with the requirements of Legislative Decree 85/16 and standard CEI EN 60079-0.

- Equipped with rechargeable batteries allowing a total autonomy of at least eight hours of continuous operation. It must be possible to check the charging status of the batteries.

### Construction characteristics

- Construction features common to all models:

- portable and compact;

- continuous operation;

- fitted with automatic scale change;

- provided with bag/case for transporting the instrument;

- small size, given the type of appliance;

- optical indicator clearly showing that the instrument is in operation;

- keyboard that allows the instrument to be switched on and off only after pressing the relevant keys for at least two seconds;

- bright or backlit display for viewing data in a dimly lit environment;

- the same applies to performing the electronic zero point on the p.p.m.v. measuring range;

- ability to display main malfunctions.

- Design features of portable equipment using flame ionisation technology:

- shoulder strap for use of the instrument;

- presence of a cylinder containing a gas mixture (e.g. 40 % hydrogen, the rest nitrogen) suitable for operating the detector. This cylinder must be small in size and of a weight suitable for use.

- Design features relevant to instruments without a laser pointer and not using flame ionisation technology:

- shoulder strap for use of the instrument;

- optical and acoustic signalling to indicate proximity of the LEL value, with adjustable alarm threshold;

- reduced weight, typically less than 3 kg (excluding accessories).

- Construction features relevant to laser pointer instruments:

- small enough for handheld use;

- reduced weight, typically less than 1 kg.

### Accessories

A list of accessories is provided below, by type of use.

- Accessories for instruments without a laser pointer:

- carpet probe for the pre-location of road leaks, complete with two-wheeled trolley;

- lance probe for the location of road leaks;

- sampling probe with flexible rod (checking for the presence of gas in confined areas or on aerial parts of gas installations);

- dust and hydrophobic filters that can be plugged into the probes and related spare parts;

- 230 VAC and/or 12VDC battery charger (if batteries are rechargeable);

- in the case of detectors used for the pre-location of gas leaks, these devices must be suitable for connection to a referenced geolocation system compatible with company maps, in order to be able to enter the results of the pedestrian pre-location directly on it and, if necessary, integrate them with the results of the pre-location carried out by means of on-board vehicle instrumentation.

- Accessories for laser pointer instruments:

- 230 VAC and/or 12 VDC battery charger;

- protective case for the instrument;

- bag/suitcase for carrying the instrument.

## Description and characteristics of portable LPG detectors

### Technical characteristics

Given that there is currently no LPG leak detection equipment using laser technology, reference is made to point 4.1.1, except as indicated below.

- The equipment must be supplied and certified with an LEL/LFL value of 2.1% by volume for LPG (reference gas: propane).

- The equipment supplied must be specifically calibrated with propane gas on each of the measuring ranges provided.

### Construction characteristics

Please refer to point 4.1.3.

### Accessories

Please refer to point 4.1.4, except for the additional use of the accessory indicated below.

- bell-shaped probe for the location of road leaks.

## Technical assessment

In relation to Legislative Decree 81/08, as amended and supplemented, Italgas Reti reserves the right, in its evaluation of the use of the proposed products, to consider the operational consequences deriving from the conditions of use of the equipment as prescribed by the Supplier in the "Use and Maintenance Manual".

# MARKING

The products covered by this T.S. must be provided with the markings required by the EU Directives and by the relevant applicable technical standards, considering that the affixing of the CE marking, together with any other markings that may be used, is subject to the provisions of Decision no. 768/2008/EC of 9 July 2008 on a common framework for the marketing of products, implemented by the legislative decrees transposing the applicable product directives, a non-exhaustive list of which is given under "Legal references" in this T.S.

They must therefore bear the conformity  marking, affixed in the manner provided for by the legal provisions in force, together with any other markings provided for by the same directives (e.g. "specific marking of explosion protection”  in accordance with the ATEX Directive, “marking relating to the disposal of electrical and electronic equipment"  in accordance with the WEEE Directive, etc.).

The marking must also contain the provisions of point 4.3 of standard CEI EN 60079-29-1.

# TYPE TEST

In the design phase of the product, the "Type Tests" required by the Reference Standards and/or the Supplier's design standards and by the standards referred to therein, for verifying the intended end use of the product must have been performed.

In particular, equipment with measuring ranges 0  100 % LEL and 0  100 % volume, with the exception of equipment using laser technology, must be type-tested by the Supplier in accordance with CEI EN 60079-29-1.

# DOCUMENTATION REQUIRED FROM THE SUPPLIER

## Supplier

- The potential Supplier must be qualified as an Italgas Supplier and must be able to provide after sales services directly in Italy.

## Documentation for the company approval process

The potential Supplier must submit the following documentation to the SERTEC-NORM Department at the tender stage:

- “Declaration of Conformity” with this technical specification and any specific documents mentioned in it (legislative provisions, applied standards, etc.) to which the Supplier refers for the design and construction of the product, in accordance with the requirements of UNI CEI EN ISO/IEC 17050-1 (see the example in APPENDIX 1 to this T.S.).

The following documentation must also be submitted to SERTEC-NORM and SERTEC-LAB:

- “EU declaration of conformity”, written in Italian, as per the implementing Decrees of the applicable product directives;

- copies of "EU Type Examination Certificates" where required (e.g. according to the ATEX Directive);

- copies of the brochures and technical bulletins of the products covered by this T.S.;

- a copy of the "Use and Maintenance Manual" (identical to the one supplied with the equipment and its components), written in Italian, which includes at least the information indicated in point 4.4 of standard CEI EN 60079-29-1 and point 5 of standard CEI EN 61187 and which complies with Legislative Decree 81/08 insofar as relevant;

- copy of the test report for tests carried out by the Manufacturer, stating how they were performed and the results;

- a statement certifying ownership to market the product;

- Whatever else the potential Supplier considers useful to better qualify the proposed product.

The potential Supplier must also send a sample of the device to the SERTEC-LAB unit for the tests according to point 8 of this T.S.

At the end of this process, a positive or negative result regarding approval for company use will be issued.

## Test Documentation

The Supplier must issue a “type 3.2” Inspection Certificate in accordance with UNI - EN 10204, which states that the products provided are compliant with the requirements of these T.S. which provides the test results.

Registration of the full production/test documentation completed by the Supplier must be kept for ten years under its responsibility.

## Supply Documentation

Each individual supply of materials covered by this T.S. must be accompanied by:

- "EU Declaration of Conformity" , in Italian, in accordance with legislative decrees implementing the Directives applicable to the product (e.g. ATEX, EMC, etc.);

- “EU type examination certificate" (Directive 2014/34/EU);

- copy of the "Installation, Use and Maintenance Manual" (in Italian) for the correct use of the product;

- documents compliant with Legislative Decree no. 286 of 21 November 2005, as amended.

# TESTS AND INSPECTIONS

For each portable detector model, the SERTEC-LAB unit will perform tests and checks to verify compliance with the requirements of this T.S. on a test sample.

The Supplier must ensure that the calibration benches supplied to the SERTEC-LAB unit are updated whenever hardware/software changes are made to the products or benches. Any modifications must be approved by Italgas Reti in advance.

## Test sample

The sample for these tests will be as follows:

- Detectors not equipped with laser technology sensors:

- two methane gas detectors for each model to be tested and, if the detector is not equipped with a dual function for selecting the measurement of methane gas and LPG, also an LPG detector.

- Detectors fitted with laser technology sensors:

- a detector.

## Type of tests

The SERTEC-LAB unit will check the following:

- presence and conformity of the marking label;

- completeness and clarity of the user and maintenance manual included with the instrument (in cooperation with SERTEC-NORM);

- technical characteristics;

- construction characteristics;

- functionality of the detector;

- functionality of the accessories;

- functionality of calibration benches (except for detectors equipped with a laser pointer).

## Right of the client

In order to verify the conformity of the proposed products with the contents of this technical specification and the cited standards to which reference is made, it is Italgas Reti's right, for each purchase order, to request from the Supplier:

- the production of further relevant technical and/or graphic documentation with the indicated characteristics;

- the performance of any additional laboratory and/or field tests in relation to the required product characteristics.

# CONDITIONS OF SUPPLY

Supplier's Standard in compliance with Legislative Decree no. 152 of 3 April 2006 as amended and supplemented.

# REFERENCE STANDARDS

- Law no. 186 of 1 March 1968 - "Provisions concerning the production of materials, equipment, machinery, electrical and electronic installations and systems";

- Legislative Decree No. 152 of 03 April 2006 as subsequently amended and supplemented “Environmental standards”;

- Legislative Decree no. 194 of 06 November 2007, as amended. - "Implementation of Directive 2004/108/EC on electromagnetic compatibility, and Directive 2014/30/EU of 26 February 2014 on the harmonisation of the laws of the Member States relating to electromagnetic compatibility (recast) providing for its repeal";

- Legislative Decree no. 86 of 19 May 2016 - "Implementation of Directive 2014/35/EU on the harmonisation of the laws of Member States relating to the making available on the market of electrical equipment designed for use within certain voltage limits";

- Legislative Decree no. 85 of 19 May 2016 – "Implementation of Directive 2014/34/EU on the harmonisation of the laws of the Member States relating to equipment and protective systems intended for use in potentially explosive atmospheres";

- Legislative Decree no. 49 of 14 March 2014, as amended. - 'Implementation of Directive 2012/19/EU on waste electrical and electronic equipment (WEEE)';

- Legislative Decree No. 81 of 09 April 2008 as subsequently amended and supplemented - “Implementation of Article 1 of Law no. 123 of 3 August 2007 on the protection of health and safety in the workplace”;

- Legislative Decree No. 286 of 21 November 2005 as subsequently amended and supplemented - “Provisions for the regulatory reorganisation of the regulated liberalisation of the road haulage business”;

- CEI EN 60529 - “Degrees of protection provided by enclosures (IP Code)”;

- CEI EN 61187 - "Electrical and electronic measuring equipment - Documentation";

- UNI CEI EN ISO/IEC 17050-1 - "Conformity assessment - Supplier's declaration of conformity - Part 1: General requirements”;

- CEI EN 60079-0 - Explosive atmospheres - Part 0: Equipment - General Requirements";

- CEI EN 60079-29-1 – "Explosive atmospheres - Part 29 -1 Flammable gas detectors – General and performance requirements";

- CEI EN 61010-1 – "Safety requirements for electrical equipment for measurement, control and laboratory use - Part 1: General Requirements";

- CEI EN 60825-1 – "Laser product safety - Part 1: Equipment classification and requirements";

- CIG Guideline no. 16 - Performing scheduled inspections and locating leaks in the distribution network for gases with density ≤ 0.8 and with density > 0.8;

- UNI EN 1839 – "Determination of explosion limits of gases and vapours".

# APPENDICES

- APPENDIX 1 -	“Sample declaration of conformity with this technical specification”

## APPENDIX 1

Supplier's Declaration of Conformity

(in accordance with the UNI CEI EN lSO/IEC 17050-1 standard)

(1)	No. ...............................................

(2)	Name of the issuer: ..............................................................................................................................................................................

Address of the issuer: .........................................................................................................................................................................

(3)	Subject of the declaration: ...................................................................................................................................................................

...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

The subject of the above declaration complies with the requirements of the following documents:

Document nos.	Title	Edition/Date of issue

(4)	S.T.V.F.C. 9993302	Portable methane gas and LPG detectors	Edition 2 of 15/06/2017

Additional information:

(5) 	...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

...................................................................................................................................................................................................................

Signed for and on behalf of:

.............................................................

.............................................................

(Place and date of issue)

(6) 	............................................................. 	......................................................................................................................

(Name and position)				(Signature or equivalent mark authorised by the issuer)

KEY:

- The Declaration must be issued on headed paper or, if visible, on applied stamps and be uniquely identified (e.g. by a log number).

- The person issuing the Declaration (Supplier) must be unequivocally specified.

- List all the products for which approval for company use is requested (name, type, model), stating any codes assigned by the Supplier itself and, where existing, the relevant material codes assigned by Italgas.

- Include references to the Italgas Reti documents for the proposed products.

- Indicate any deviations from the reference documentation and any other information deemed useful by the Supplier for product evaluation purposes (see note 6) of appendix A of UNI CEI EN ISO/IEC 17050-1).

- Include the handwritten signature of the person authorised to issue the Declaration, stating their full Name, Surname and Position within the Organisation.

Supplier means the party responsible for the design and construction of the products or its legal representative.

Lower flammability limit (LFL): percentage by volume of flammable gas or vapour in the air below which an explosive gas atmosphere is not formed. The definition of the Lower Flammability Limit and the Lower Explosive Limit (LEL) are considered synonymous in CEI EN 60079-29-1 and UNI EN 1839.